- ELECTROTEK ANAEROBIC WORKSTATION

- AW4500 SINGLE / DUAL GAS TOUCHSCREEN CONTROLLER

- CONTENTS

| 1. | OVERVIEW | 1 |
| --- | --- | --- |
| 2. | GENERAL FEATURES |  |
|  | a) Startup Page | 2 |
|  | b) Main Page | 2 |
|  | c) Temperature Controller Page | 4 |
|  | d) Humidity Controller Page | 5 |
|  | e) O2 Display Page (if fitted) | 6 |
|  | f) Main Settings Page | 7 |
|  | g) Advanced Settings Page | 8 |
|  | h) Alarms List Page | 10 |

1. OVERVIEW

This document refers to the touchscreen control panel used in Electrotek’s AW400 and AW500 anaerobic workstations, both single and dual gas types. It should be used in conjunction with the operator’s manual.

2. GENERAL FEATURES

a) 	Startup

On applying power to the system, an intro screen will be displayed for five seconds. This 		contains two pieces of information – the serial number and model number of the cabinet.

After five seconds, the main control page will be displayed and the pulsed alarm will sound 		to warn the user that the mains supply has been interrupted.

As with most audible alarms on the cabinet, this alarm can be silenced and in this case 			cancelled by the silence alarm button. In most cases, the silence alarm button will only 			cancel the audible alarm and will not cancel the display of the alarm condition. If the cabinet 		loses its mains supply, the battery operated continuous alarm will sound until the mains is 		restored.

b)	Main Page

Left buttons

Start Lock Flush – 	Initiates lock flushing, provided that both doors are closed and there 					is sufficient gas pressure.

Silence Alarm - 	Silences the audible pulsed alarm. In the case of a gas leak or mains 					failure warning, the button also clears the fault. In all other cases, the 					fault warning remains until the fault is cleared.

Comm Switch - 	The commissioning switch allows the lock flush to continue past the 					five minute limit for commissioning purposes. While activated, the 						yellow warning bar will announce this.

Cabinet Lights - 	Switches the main cabinet lights on and off.

Spot Lights - 		Switches the spot light(s) on and off.

Indicator Pump - 	Switches the socket controlling the indicator pump on and off.

Indicator Lamps

Lock Flushing - 	Lit whenever the cabinet is flushing the lock.

Outer Door Open - 	Lit whenever outer door is open.

Inner Door Open - 	Lit whenever inner door is open.

Heater - 		Lit whenever the heater element is powered. This will flash when the 					cabinet is on or close to its set temperature.

Dehum Fan - 		Lit whenever the dehumidifier fan is active.

Temp Alarm - 		Lit whenever the temperature of the cabinet is outside the constraints 					set by the temperature high and low alarms (when enabled).

Gas Leak - 		Lit whenever the system has detected a gas leak, i.e. the mixed gas 					solenoid has been active for a continuous period of more than two 					minutes.

Hum Alarm - 		Lit whenever the humidity of the cabinet is outside the constraints 						set by the humidity high and low alarms (when enabled).

O2 Alarm - 		Lit whenever the oxygen level of the cabinet is outside the 							constraints set by the oxygen level high and low alarms (when 						enabled). (Only visible if sensor unit fitted)

Mixed Gas Flow - 	Lit whenever the mixed gas solenoid is active.

N2 Gas Flow - 		Lit whenever the N2 gas solenoid is active.

Low MG Pressure - 	Lit whenever the mixed gas pressure is below that set on the mixed 					gas pressure switch.

Low N2 Pressure - 	Lit whenever the N2 gas pressure is below that set on the N2 gas 						pressure switch.

Centre Displays

O2 %			Displays the current O2 concentration. (Only visible if sensor unit 						fitted)

T °C			The larger digits display the current cabinet temperature. The smaller 					digits display the set temperature.

RH%			The larger digits display the current cabinet humidity. The smaller 						digits display the set humidity.

Info Bars

Top Info Bar - 		Displays alarms and information such as commissioning mode or 						heater system autotune. Only visible under alarm or necessary 						information conditions.

Centre Info Bar - 	Displays the current aerobic / anaerobic condition of the cabinet.

Lower Info Bar - 	Displays the time remaining on a lock flush. Only visible during lock 					flushes.

Lower Right Buttons

O2 PB - 		Press to access the O2 settings page. (Only visible if sensor unit 						fitted)

Settings PB - 		Press to access the settings / advanced settings pages. If a 						password has been set, this will be required for access.

Humidity PB - 		Press to access the humidity controller page.

Temperature PB - 	Press to access the temperature controller page.

c)	Temperature Page

Graph - 		Shows the temperature setting and actual temperature data collected 					since the last power cycle of the cabinet. Pressing the ‘M’ button on 					the left allows viewing of the recorded data, while pressing the ‘M’ 						button again returns the display to current data. While in memory 						mode, a pair of arrow buttons at the bottom of the graph allow the 						user to move through the data in terms of time. Date and time are 						displayed on the left and right of the X-axis.

Pressing the ‘G’ button applies a horizontal grid to the display.

Actual - 		Displays the current cabinet temperature.

Setting - 		Displays the desired set point of the temperature. To adjust this, 						press the grey setting display and a numerical keypad will appear, 					as below.

High / Low Alarm - 	These display the high and low limits of temperature for the 						temperature alarm. They can again be adjusted by pressing the grey 					displays and entering a new value on the numerical keypad.

Heater Indicator - 	Lit whenever the heater element is powered. This will flash when the 					cabinet is on or close to its set temperature.

Heater Enable - 	Switches the heating system on and off.

Alarm Enable - 	Switches the temperature alarm system on and off. It is 							recommended that this is switched off when setting up the cabinet.

Back Button - 		(Right arrow) This button navigates back to the main page.

d)	Humidity Page

Graph - 		Shows the humidity setting and actual humidity data collected 						since the last power cycle of the cabinet. Pressing the ‘M’ button on 					the left allows viewing of the recorded data, while pressing the ‘M’ 						button again returns the display to current data. While in memory 						mode, a pair of arrow buttons at the bottom of the graph allow the 						user to move through the data in terms of time. Date and time are 						displayed on the left and right of the X-axis.

Pressing the ‘G’ button applies a horizontal grid to the display.

Actual - 		Displays the current cabinet humidity.

Setting - 		Displays the desired set point of the humidity. To adjust this, 						press the grey setting display and a numerical keypad will appear, 					as below.

High / Low Alarm - 	These display the high and low limits of humidity for the 							humidity alarm. They can again be adjusted by pressing the grey 						displays and entering a new value on the numerical keypad.

Fan Indicator - 	Lit whenever the dehumidifier fan is powered.

Fan Enable - 		Switches the dehumidifier system on and off.

Alarm Enable - 	Switches the humidity alarm system on and off. It is 							recommended that this is switched off when setting up the cabinet.

Back Button - 		(Right arrow) This button navigates back to the main page.

e)	O2 Page

(This page is only available if the O2 sensor unit is fitted)

Graph - 		Shows the actual O2 data collected since the last power cycle of the 					cabinet. Pressing the ‘M’ button on the left allows viewing of the 						recorded data, while pressing the ‘M’ button again returns the display 					to current data. While in memory mode, a pair of arrow buttons at the 					bottom of the graph allow the user to move through the data in terms 					of time. Date and time are displayed on the left and right of the X-						axis. Pressing the ‘G’ button applies a horizontal grid to the display.

Actual - 		Displays the current cabinet humidity.

High / Low Alarm - 	These display the high and low limits of humidity for the 							humidity alarm. To adjust this, press the grey setting display and a 					numerical keypad will appear, as below.

Alarm Enable - 	Switches the O2 alarm system on and off. It is recommended that this 					is switched off when setting up the cabinet.

Back Button - 		(Right arrow) This button navigates back to the main page.

f) 	Main Settings Page

Network Address - 	This is the network address of the cabinet. Press each segment of 					the address to set / edit the address.

Subnet Mask - 	This is the subnet mask of the cabinet. Again, these are set by 						individual segment.

Default Gateway - 	The default gateway of the connection or network. Again, these are 					set by individual segment.

All changes to these settings must be confirmed by pressing the ‘Apply’ button 			when complete.

Unit Identifier - 	This is the cabinet name. It is used to identify the cabinet to the data 					logging or remote control software. To enter / edit, press the grey 						area and enter the name on the alphanumeric keypad which will 						appear as below. This is only saved on a power cycle, so the unit 					must be restarted.

System Time / Date - 	This can be set manually or by the datalogging software. To set 						manually, press the grey areas and input the details via the numeric 					keypad which will appear.

Both Doors -		This enables the optional alarm which sounds if both doors are 			Open Alarm		opened at the same time.

Autotune Temp - 	This starts the temperature controller autotune. This is an automatic 					procedure which trims the temperature controller software for its 						particular environment. The procedure can take around an hour to 					complete and will cause temperature overshoot, so it is not advisable 					to perform this with experimental material in the cabinet. This 						procedure has already been done during manufacturing and should 					not need to be repeated on site.

Alarms List - 		This opens the alarms list page.

Advanced Settings - 	This opens the advanced settings page.

Back Button - 		(Right arrow) This button navigates back to the main page.

g)	Advanced Settings Page

O2 Offset - 		This allows an offset of +/-5% to be entered to correct for O2 sensor 					variation and sensor positioning inside the cabinet during calibration. 					To enter an offset, press the grey area and enter the offset using the 					numerical keypad which appears. (Only available if sensor unit 						fitted)

Humidity Offset - 	This allows an offset of +/-10% to be entered to correct for humidity					sensor	variation and sensor positioning inside the cabinet during 						calibration. To enter an offset, press the grey area and enter the 						offset using the numerical keypad which appears.

Temperature Offset - 	This allows an offset of +/-5°C to be entered to correct for 							temperature sensor variation and sensor positioning inside the 						cabinet during calibration. To enter an offset, press the grey area and 					enter the offset using the numerical keypad which appears.

Cabinet Serial no. -	This is the cabinet serial number which appears on the intro screen. 					It is useful when ordering spares and reporting any issues to 						engineers. This should not be changed by the user.

Set New Engineer - 	This button allows the user to enter change the password which 			Password		protects the main and advanced settings pages. If no password is set 					or it is set to blank and saved, a user will be able to access theses 					settings pages by pressing the settings button on the man page. 						Once a password has been entered and saved, this will prevent 						unauthorised access. When pressed, this button will bring up a 						password entry dialog box as below.

Pressing the grey area will access the alphanumeric keypad. Once 					the password is set, press the back button to return to the previous 					page.

Select Cabinet - 	Pressing one of these buttons selects the model of cabinet being 						used. This will be set during manufacture and should not be changed 					by the user.

Back Button - 		(Right arrow) This button navigates back to the main page.

h) 	Alarms List Page

The alarms list keeps a record of the last fifty alarm conditions. These records are 			unaffected by power cycling. Each record consists of a reference number, date and time 			stamp, and details of the alarm condition.

Left buttons

Go To First Record -	Sets the top row to alarm record no.1.

(Top button)

Go Up 5 - 		Moves the top row up five places.

(Second button down)

Go Down 5 - 		Moves the top row down five places.

(Third button down)

Go To Last Record - 	Sets the bottom row to alarm record no.50.

(Bottom button)

Others

Clear Alarm List - 	Clears the whole list. If pressed, a confirmation screen will appear as 					below.

Back Button - 		(Right arrow) This button navigates back to the main page.